Radiation Protection					Name:____________________________

Score:

LAB #2: Radiation Measuring Devices

Materials needed: Several self-reading pocket dosimeters with a scale of 1-200 mR or 1-1000

mR; Battery-powered charger for pocket dosimeters; Ion chamber; Geiger

counter. Lead aprons or mobile barriers. Lead gloves.

PART A: Pocket Dosimeter Vs Ion Chamber

Read through all instructions before proceeding:

Each student or group is assigned one pocket dosimeters with a scale of 1-200 mR or 1-1000 mR. Holding it down on the charger, turn the knob to adjust the needle to the lower portion of the scale. It does not have to be on zero. Be sure it is not to the left of zero. If the needle is high on the scale after the first exposure is made, you may have to re-set it lower for the second exposure.

Be sure not to jolt the pocket dosimeters as you place them on the table or remove them.

If you are in a group, for all readings double-check each other to make sure you agree. Note the initial reading on the dosimeter before making any exposures and write it below. To find the actual exposure to fill in below, you will need to subtract this amount from your “after-exposure” reading:

Reading the Ion Chamber:  The ion chamber must be read immediately after the exposure for

accuracy. If it reads in R, move the decimal 3 places to the right to get the reading

in mR.

Exposure:  On the x-ray machine, set 66 kVp. Use 10 mAs for exposure #1 and 20 mAs for

exposure #2. Keep the kVp, collimation, distance and all other factors equal.

# Open the collimated field wide enough to fit all the devices in one exposure. Place

the lab ionization chamber in the center, set for “dose” or “integrate.” Also carefully place all of the pocket dosimeters well within the light field. Record your results below.

Preliminary Pocket Dosimeter Readings:

1st Initial Reading =                    mR		After Exposure Reading =                     mR

2nd  Initial Reading =                    mR		After Exposure Reading =                     mR

Results:

First mAs . . . . . . .  =	                         	Second mAs . . . .   = 	______________

First Ion 					Second Ion

Chamber Reading  =                           	Chamber Reading  =   ______________

First Pocket					Second Pocket

Dosimeter Reading =                          	Dosimeter Reading =  ______________

Questions:

1. According to the ion chamber, how close did a doubling of the mAs come to actually

producing twice the radiation? [The instructor can help you convert this into a

percentage of inaccuracy if you like.] Was the 2nd exposure too high or too low?

2. According to your pocket dosimeter, how close did a doubling of the mAs come to actually

producing twice the radiation? Was the 2nd exposure too high or too low?

3. Compare the First pocket dosimeter reading with the First ion chamber reading. Generally, the ion chamber is more reliable. How closely did your team’s pocket dosimeter agree with the ion chamber? Did the pocket dosimeter tend to read higher or lower than the ion chamber? Does the same tendency show up when comparing the Second set of readings?

PART B (Group Lab):  mA Vs. Dose Rate

In the lab, set the ion chamber for “dose rate.” In small groups, use shielding (lead aprons or mobile barriers), and observe the meter reading up close while the instructor makes two exposures.

The instructor will set the second mA to double the first mA. The meter will fluctuate and you will only have a few seconds to decide where it averages out. Note your readings below.

Recommended techniques:

Large FS, 60 kVp, 4 seconds: 25 mA for 1st exposure, 50 mA for 2nd exposure.

First mA  . . . . . . .  =	                         	Second mA  . . . .   =   _____________

First Dose Rate 				Second Dose Rate

Chamber Reading  =                           	Chamber Reading  =   _____________

1. Did your readings of dose rate go up roughly proportional to the increase in mA? Was the

second reading too high or too low for a doubling?

2. mA is defined as the                         of flow of electricity through the x-ray tube.

3. Here we checked the actual dose rate produced by the mA stations. Of the three modes of

operation for a detection device (pulse, dose rate, and integrated dose), which one would be most appropriate for checking the accuracy of total mAs?

PART C: Appropriateness of Geiger Counter for Checking X-Ray Machines

Place the Geiger counter in the x-ray beam while you makes a long exposure. NOTE: If a button must be held down on the Geiger counter to operate it, it is recommended that the instructor hold the device in the beam while wearing lead gloves. Be careful not to get the lead gloves over the detection chamber area.

Listen carefully for any clicks from the Geiger counter during or after the exposure.

Recommended technique:

Large FS, 60 kVp, 50 mA, 4 seconds.)

1. What was the response, if any, of this device to the x-ray machine during the exposure?

Were there any clicks? Was there continuous clicking? Describe what you heard:

2. Which one of the four characteristics of radiation measuring devices: (sensitivity, accuracy, resolution time or range) is directly related to this result?

3. Explain this result in terms of the characteristic you identified in #3 above:

4. If you heard an unusual number of clicks after the exposure ended, what phenomenon from the

voltage response curve would account for this?