Quality Control						Name:_____________________________

Score:

Lab #3 Half Value Layer (HVL) as a Function of Added Filtration

-By Bob Grossman

Materials needed: Tissue-equivalent abdomen or lumbar spine or hip phantom, 3” of paper or 3 blocks of wood, or 3 packs of hand sanitizer, or 3 bags of saline.

Procedure

Place the phantom in the center of the x-ray table and set a technique of 90 kVp.  Be sure to turn on the AEC. If you are using the abdomen or a lumbar spine phantom, verify that only the center detector cell is selected.   If you are using a hip phantom, verify that only the center right(or left) detector cell is selected. Verify that the density control is set to 0.

For each addtional filtration level, take one of each of the following: an exposure with no additional attenuators above your phantom, 1 block of wood (or hand sanitizer packs, or saline bags,  or inches of paper), 2 attenuators, and 3 attenuators.

Added Filtration Condition 2 should be the most additional filtration your system supports, and Added Filtration Condition 1 is more filtration than no additional filtration but less filtration than Condition 2.

For each exposure, record the applied mAs and the exposure indicator value.

| Exposure | Added Filtration | Additional Attenuators | Applied mAs | Exposure Indicator |
| --- | --- | --- | --- | --- |
| 1 | None | None |  |  |
| 2 |  | 1 |  |  |
| 3 |  | 2 |  |  |
| 4 |  | 3 |  |  |
| 5 | Added Filtration Condition 1 | None |  |  |
| 6 |  | 1 |  |  |
| 7 |  | 2 |  |  |
| 8 |  | 3 |  |  |
| 9 | Added Filtration Condition 2 | None |  |  |
| 10 |  | 1 |  |  |
| 11 |  | 2 |  |  |
| 12 |  | 3 |  |  |

Analysis

- Did your exposure indicator stay close to optimal for all exposures?  Were there any exposures resulting in over-exposed or under-exposed exposure indicator values?

- Create a graph showing applied mAs as a a function of Attenuators.  Make the Y axis applied mAs and the X axis the number of additional attenuators (none,1, 2 & 3).  Create three series:  No added filtration, Added Filtration Condition 1, and Added Filtration Condition 2, so your graph will demonstrate three lines.

*Note if any of your exposures resulted in an error stating “applied mAs has reached the maximum value” do not include that exposure in your graph.  This condition is most likely to occur for Added Filtration Condition 2 with 3 Additional Attenuators present and will be accompanied with an exposure indicator value indicative of a technique which resulted in an under-exposed radiograph.

- Using the No Added Filtration line on your graph, find the X value when the Y value equals 2 times the applied mAs measured for Exposure 1.

No Added Filtration HVL = ___________________

*Note the units for HVL will be either blocks of wood or  packs of hand sanitizer, or bags of saline,  or inches of paper, depending on what material you used for attenuators.

- Using the Added Filtration Condition 1 line on your graph, find the X value when the Y value equals 2 times the applied mAs measured for Exposure 5.

Added Filtration Condition 1 HVL =_________

- Using the Added Filtration Condition 2 line on your graph, find the X value when the Y value equals 2 times the applied mAs measured for Exposure 9.

Added Filtration Condition 2 HVL =_________

- Using the applied mAs from Exposure 5 and Exposure 1, calculate the % reduction in intensity due to Added Filtration Condition 1.

- Using the applied mAs from Exposure 9 and Exposure 1, calculate the % reduction in intensity due to Added Filtration Condition 2.

- Describe how filtration effects intensity and HVL.

Optional Additional Analysis

Measure ROI values of a high attenuator and a low attenuator.    If your low attenuator has a mean value higher than your high attenuator, then your ROI values are reporting the level of darkness.  If your high attenator ROI mean value is higher than your low attenuator mean value, then your ROI values are already reporting the level of brightness and no conversion is required.

To convert level of darkness into brightness values on a 12-bit system convert each ROI value into a brightness value using the formula:

Brightness Value = 212  –  ROI Mean Value = 4096 – ROI Mean Value

*For a 14-bit system:  Brightness Value = 214  –  ROI Mean Value = 16384 – ROI Mean Value

Calculate contrast as the ratio of the High Attenuator Brightness Value / Low Attenuator Brightness Value.

Contrast = 	High Attenuator Brightness Value

Low Attenuator Brightness Value

Graph or describe how contrast changes with filtration.

Depending on the Additional Filtration available you may be able to measure and calculate HVL per mm of aluminum and/or HVL per mm of copper.