Quality Control						Name:_____________________________

Score:

Lab #2 Half Value Layer (HVL) as a Function of kVp

-By Bob Grosssman

Materials needed: Tissue-equivalent abdomen or lumbar spine or hip phantom, 3” of paper, or 3 blocks of wood, or 3 packs of hand sanitizerl, or 3 bags of saline.

Procedure

Place the phantom in the center of the x-ray table and set a technique of 80 kVp.  Be sure to turn on the AEC. If you are using the abdomen or a lumbar spine phantom verify that only the center detector cell is selected.   If you are using a hip phantom verify that only the center right detector cell is selected. Verify that the density control is set to 0.

For each kVp take one of  each of the following: an exposure with no additional attenuators above your phantom, 1 block of wood (or hand sanitizer packs, or saline bags,  or inches of paper), 2 attenuators, and 3 attenuators.

For each exposure record the applied mAs and the exposure indicator value.

| Exposure | kVp | Attenuators | Applied mAs | Exposure Indicator |
| --- | --- | --- | --- | --- |
| 1 | 80 | None |  |  |
| 2 |  | 1 |  |  |
| 3 |  | 2 |  |  |
| 4 |  | 3 |  |  |
| 5 | 90 | None |  |  |
| 6 |  | 1 |  |  |
| 7 |  | 2 |  |  |
| 8 |  | 3 |  |  |
| 9 | 100 | None |  |  |
| 10 |  | 1 |  |  |
| 11 |  | 2 |  |  |
| 12 |  | 3 |  |  |

Analysis

- Create a graph showing applied mAs as a a function of # of attenuators.  Make the Y axis Applied mAs and the X axis the number of attenuators (none,1, 2 & 3).  Create three series: 80kVp, 90kVp, and  100kVp so your graph will demonstrate three lines.

- Using the 80 kVp line on your graph, find the X value when the Y value equals 2 times the applied mAs measured for Exposure 1.  80 kVp HVL = ___________________

*Note the units for HVL will be either blocks of wood, or  packs of hand sanitizer, or bags of saline, or inches of paper depending on what material you used as attenuators.

- Using the 90 kVp line on your graph, find the X value when the Y value equals 2 times the applied mAs measured for Exposure 5.  90 kVp HVL = ___________________

- Using the 100 kVp line on your graph, find the X value when the Y equals  2 times the applied mAs measured for Exposure 9.  100 kVp HVL = ___________________

- Which kVp resulted in the largest HVL?

- Describe the relationship between kVp and HVL.

Optional Additional Analysis

Measure ROI values of a high attenuator and a low attenuator.    If your low attenuator has a mean value higher than your high attenuator, then your ROI values are reporting the level of darkness.  If your high attenuator ROI mean value is higher than your low attenuator mean value, then your ROI values are already reporting the level of brightness and no conversion is required.

To convert level of darkness into brightness values on a 12-bit system, you will need to convert each ROI value into a brightness value using the formula:

Brightness Value = 212  –  ROI Mean Value = 4096 – ROI Mean Value

*For a 14-bit system:  Brightness Value = 214  –  ROI Mean Value = 16384 – ROI Mean Value

Calculate contrast as the ratio of the High Attenuator Brightness Value / Low Attenuator Brightness Value.

Contrast = 	High Attenuator Brightness Value

Low Attenuator Brightness Value

Graph or describe how contrast changes with kVp.