Radiation Protection					Name:_____________________________

Score:

LAB #3: Resolving Time and Voltage Response

Materials needed: Ion chamber with variable voltage setting; Samples of two different radionuclides with steady, high activity (recommended: cobalt 60 and cesium 137). Small samples embedded in 1-inch disks which can be stacked or otherwise placed close together are recommended.

PART A: Resolving Time

DIRECTIONS:

If the ion chamber has a re-set button, press it between each measurement. Place sample A (Cobalt 60) at a fixed, short distance (an inch or two) from the detector element. Set the counting time to 10 seconds. (Rate mode can be used, but is more difficult.) For accuracy, take 3 measurements at this time setting, and average them by dividing their sum by 3. Record this number below as RA.

Repeat the above procedure using sample B (Cesium 137) and record the average as RB.

Repeat the above procedure with both samples stacked together or otherwise placed as close together as possible. Again, average 3 measurements. Record this number as RA + RB below:

Measurements:		        RA   =   ____________

RB   =   ____________

RA + RB  =   ____________

Insert the measurements into the formula below to determine the resolving time for this ion chamber:

T   =      RA + RB  -  (RA + RB)   	=	_____________

2(RA X RB)

The resulting number will be negative. Ignore the negative sign. The resolving time is the absolute value of this number - Report it below:

Resolving Time for this ion chamber   =       ______________________________

PART B: Voltage Response Curve

DIRECTIONS: An ion chamber with variable voltage settings is required for this lab.

Place both samples of radionuclides (cobalt 60 and cesium 137) together at a short, fixed distance in front of the ion chamber detector element. Set the counting time for 10 seconds. Beginning at 500 volts (or low enough voltage that the measurement reads zero), take a measurement at increasing increments of 100 volts, and record below:

Measurements:		500 V   =     .                    .

600 V   =

700 V   =

800 V   =

900 V   =

1000 V   =

1100 V   =

1200 V   =

Plot and draw in below (on the next page) the voltage response curve for the this ion chamber device.

Label the Recombination region, the simple ionization plateau, and the proportional region

portions of the curve. (Note that the Geiger-Muller plateau is only reached operating at much higher voltages than typical ion chambers use.)