Quality Control		 			Name:_____________________________

Score:

LAB #3: Exposure Reproducibility Test

Materials needed: Ion chamber

Technique for all exposures:  70 kVp and 100 mA at 0.5 seconds

Procedure

Take five exposures using this identical technique. (For accuracy, reproducibility tests are usually done using 10 identical exposures.) Take readings immediately after each exposure, and re-set the ion chamber. Record the ion chamber read-outs below. If the meter reads in R, convert the measurements into mR by moving the decimal 3 places to the right:

Exposure #1:                mR

Exposure #2:                mR

Exposure #3:                mR

Exposure #4:                mR

Exposure #5:                mR

Analysis

1. Add up the readings for all 5 exposures and divide this sum by 5 to obtain the average

exposure:

Average Exposure   =                       mR

2. For each exposure, take the measured reading and the average from Question #1, and divide

the smaller number by the larger, then multiply by 100 to obtain the percentage of

deviation:

Exposure #1:               (Lesser mR)    =                     X   100   =                    %

(Larger mR)

Exposure #2:               (Lesser mR)    =                     X   100   =                    %

(Larger mR)

Exposure #3:               (Lesser mR)    =                     X   100   =                    %

(Larger mR)

Exposure #4:               (Lesser mR)    =                     X   100   =                    %

(Larger mR)

Exposure #5:               (Lesser mR)    =                     X   100   =                    %

(Larger mR)

3. What is the acceptable limit of deviation for exposure reproducibility? (See textbook.)

4. For acceptable reproducibility, all 5 (or 10) exposures taken at the identical technique must

fall within limits. Do all of the percentages of deviation found in Question #2 fall within

this range? List those which do not: