Computerized Radiography (CR) 			Name:_____________________________

Score:

LAB #9: Attempts to Induce Histogram Analysis Failure

Materials needed: Tissue-equivalent elbow (or similar) phantom.

Procedure

Take all exposures tabletop. Be sure to turn off the AEC and use “manual technique.”

Consult a technique chart or your instructor for a technique, or use:

Suggested test technique:   66 kVp at 4 mAs

A. On a single 14" X 17" CR plate, take a single exposure of an elbow (or similar) phantom, placing the object as far from the center as the cassette will allow and opening the field to cover the entire plate.

B. On a single 14" X 17" CR plate, take a single exposure of an elbow (or similar) phantom, centered to the cassette, but over-collimated to a field of only 4 inches square.

C. If two knee (or similar) phantoms are available, on a single 14" X 17" CR plate, take a single

exposure simulating a bilateral projection of both phantoms. Separate the phantoms well, with the central ray in the space between them and the field filling the cassette.

Process all exposures. Print them out only with the instructor’s permission.

Analysis

1. Do any of the images demonstrate unexpected levels of brightness, contrast or gray scale?

For Exposure A:	yes / no

For Exposure B:	yes / no

For exposure C:	yes / no

2. Do any of the images demonstrate unusual mottle, aliasing, fog or other artifacts?

For Exposure A:	yes / no

For Exposure B:	yes / no

For exposure C:	yes / no

3. If you answered “yes” to any of the above, explain why (see textbook):