Quality Control		 			Name:_____________________________

Score:

LAB #4: Focal Spot Size Test and Blooming (Demonstration*)

*NOTE: This test is designed to be performed with x-ray film in a “direct-exposure” holder, so that crystal size and phosphor layer thickness in screen or CR cassettes are not contributing to blur and thus skewing the measurement of the focal spot size. However, it can be done with a CR system in order to merely demonstrate the process to students, and is included here for that purpose, but it should be emphasized that these resulting measurements are not an accurate indication of the actual focal spot size.

This lab can be used to provide an excellent demonstration of the effects of mA on focal spot blooming, since it is relative to the original focal spot measurement.

There are three methods of measuring focal spot condition - the pinhole camera, the line-group test (Wisconsin) tool method, and the “star” type focal spot/resolution test template method. The first is not recommended because of the danger of over-loading the x-ray tube with the very high techniques required. The second is much less accurate, and requires an accompanying numerical table for conversion of line-pairs per millimeter (LP/mm) into focal spot dimensions. Since the star pattern method is simple to perform, and also much more accurate, only this method is presented here.

Materials needed: “Star” type focal spot/resolution test template; Ruler with mm measurements.

Procedure

Place the star test template at distance precisely one-half of the SID. (Measure the SID accurately by using a tape measure and measuring from the “X” or “+” mark at the anode end of the x-ray tube housing.) If the test template is supplied with a holder stand (usually 14 inches tall), set the SID exactly to double this amount (28 inches for the 14" stand). If the pattern of cut lines in the template forms a “Maltese cross,” (instead of radiating in all directions), it is important to position this cross pattern so that its branches run perpendicular and parallel to the table rather than diagonally - The cross-hairs of the light field should run down the middle of the line patterns.

(If a template is used that radiates lines in all directions, you must use markers to indicate the directions of the cathode and anode ends of the table, because the focal spot dimensions must be measured parallel and perpendicular to this axis and not diagonally.)

Make three exposures with the techniques listed below. The kVp is varied only to compensate for changes in mA in order to maintain overall exposure level. For a particular laboratory x-ray machine, in the event that focal spots cannot be selected independently of the mA stations, alternate techniques are suggested below. Label the exposures with lead markers to keep track of them.

Exposure #1: Small focal spot, 50 mA, 0.1 second, 60 kVp

Exposure #2: Large focal spot, 50 mA, 0.1 second, 60 kVp [or 300 mA, 0.016 second]

Exposure #3: Large focal spot, 300 mA, 0.1 second, 60 kVp

Analysis

1. On the resulting star pattern images, for each set of dimensions (lengthwise and crosswise to

the anode-cathode axis of the x-ray tube), scan visually from the periphery of the pattern

inward until you find the first point of blur where the lines become indistinct from each

other. (Apparent recurrences of the lines further inward are caused by false resolution in

which the lines are blurred so much that they begin to overlap each other.) With pencil or

permanent ink, mark a line across this point on each arm of the star pattern, so there are 4

marked lines in total.

2. With a metric ruler, measure the diameter between each pair of the marks you made - this is

the distance through the center of the pattern between the two lines. Record these

measurements below:

Exposure #1: Small focal spot:	Lengthwise:                mm

Crosswise:                 mm

Exposure #2: Large focal spot:	Lengthwise:                mm

Crosswise:                 mm

Exposure #3: Large FS, High mA:	Lengthwise:                mm

Crosswise:                 mm

3. Multiply each of these blur diameters by the factor 0.35 to obtain the focal spot dimensions,

(or other factor indicated by the manufacturer of the test template.) Record these

dimensions below:

Exposure #1: Small focal spot:	Lengthwise:                mm

Crosswise:                 mm

Exposure #2: Large focal spot:	Lengthwise:                mm

Crosswise:                 mm

Exposure #3: Large FS, High mA:	Lengthwise:                mm

Crosswise:                 mm

4. Compare the focal spot dimensions for the small focal spot (Exposure #1) to those of the large

focal spot (Exposure #2) using the same technique; How much larger is the large FS?

(You can calculate the ratio of increase by simply dividing the larger number by the

smaller number. For example, 1.83 would mean that it is 1.83 times larger.) Is the large

focal spot close to double the size of the small focal spot?

5. For the large focal spot, compare the focal spot dimensions between the low mA setting of 50

mA (Exposure #2) and the high mA setting of 300 mA (Exposure #3); Have the

dimensions changed at the higher mA? If so, which way and by how much? Why? (See

textbook.)