Demonstrating Basic Digital Concepts			Name:_____________________________

Score:

Lab #1 Automatic Exposure Control

-By Bob Grossman

Materials needed: Tissue-equivalent abdomen or lumbar spine or hip phantom.

Procedure

Part A: Applied mAs as a function of positioning accuracy

Place the phantom in the center of the x-ray table and set a technique of 80 kVp. Be sure to turn on the AEC.    If you are using the abdomen or a lumbar spine phantom, verify that only the center detector cell is selected.   If you are using a hip phantom, verify that only the center right detector cell is selected. Verify that the density control is set to 0.

For each exposure, record the applied mAs and the exposure indicator value.

| Exposure | Condition | Applied mAs | Exposure Indicator |
| --- | --- | --- | --- |
| A1 | Phantom centered |  |  |
| A2 | Phantom off centered to the right 1” |  |  |
| A3 | Phantom off centered to the left 1” |  |  |
| A4 | Phantom off centered superiorly 1” |  |  |
| A5 | Phantom off centered inferiorly 1” |  |  |

Part B: Applied mAs as a function of the density control

Place the phantom in the center of the x-ray table and set a technique of 80 kVp. Be sure to turn on the AEC.    If you are using the abdomen or a lumbar spine phantom, verify that only the center detector cell is selected.   If you are using a hip phantom, verify that only the center right detector cell is selected. Verify that the density control is set to 0.

For each exposure, record the applied mAs and the exposure indicator value.

| Exposure | Condition | Applied mAs | Exposure Indicator |
| --- | --- | --- | --- |
| B1 | Density control set to 0 *Note you can use Exposure A1. |  |  |
| B2 | Density control set to -3 |  |  |
| B3 | Density control set to -2 |  |  |
| B4 | Density control set to -1 |  |  |
| B5 | Density control set to +1 |  |  |
| B6 | Density control set to +2 |  |  |
| B7 | Density control set to +3 |  |  |

Analysis: Part A

- Which condition resulted in the most optimal exposure indicator value?

- Using the exposure indicator value, determine which condition resulted in the most over-exposure.

- Using the exposure indicator value, determine which condition resulted in the most under-exposure. Is this the condition which resulted in the lowest applied mAs?  Is mottle present in this radiograph?

- How did the applied mAs and exposure indicator change when you moved your centering to the right?

- How did the applied mAs and exposure indicator change when you moved your centering to the left?

- How did the applied mAs and exposure indicator change when you moved your centering superiorly?

- How did the applied mAs and exposure indicator change when you moved your centering inferiorly?

- Which position resulted in the highest applied mAs?  Why do you think this position resulted in the highest applied mAs?

- Which position resulted in the lowest applied mAs?  Why do you think this position resulted in the lowest applied mAs?

Analysis: Part B

- Which density control setting resulted in the most optimal exposure indicator value?

- Create a graph of applied mAs as a function of the density control setting.  Make  your Y axis applied mAs and your X axis the density control setting.

- By what % does the applied mAs change for each change of the density control?

- Under what conditions is it appropriate to use the density control?

- Compare exposure 2 (density control = -3) with exposure 7 (density control = +3) in terms of brightness and contrast.

- Which density control setting resulted in the most over-exposed radiograph?  What was the exposure indicator value?

- Which density control setting resulted in the most under-exposed radiograph?  Is mottle present?  What was the exposure indicator value?

- Describe the relationship between applied mAs and the density control.

Optional Additional Part B Analysis

Convert Exposure Indicator values into Exposure values as expressed in uGy.  Graph or describe Exposure Indicator Value (or Exposure) as a function of density control.

Measure ROI values of a high attenuator and a low attenuator.    If your low attenuator has a mean value higher than your high attenuator, then your ROI values are reporting the level of darkness.  If your high attenuator ROI mean value is higher, than your low attenuator mean value, then your ROI values are already reporting the level of brightness and no conversion is required.

To convert level of darkness into brightness values on a 12-bit system, you will need to convert each ROI value into a brightness value using this formula:

Brightness Value = 212  –  ROI Mean Value = 4096 – ROI Mean Value

*For a 14-bit system:  Brightness Value = 214  –  ROI Mean Value = 16384 – ROI Mean Value

Calculate contrast as the ratio of the High Attenuator Brightness Value / Low Attenuator Brightness Value.

Contrast = 	High Attenuator Brightness Value

Low Attenuator Brightness Value

Measure ROIs for a high attenuator, a low attenuator and backgroud.   If your ROI values are reporting level of darkness, then calculate the brightness values as shown above.

Graph or describe brightness values for the three ROIs (High Attenuator, Low Attenuator, and Background) as a function of applied mAs.

Calculate Signal / Noise Ratio (SNR).

SNR = High Attenuatory Brightness Value / Background Brightness Value

Graph or describe SNR as a function of Applied mAs.