Radiation Protection					Name:_____________________________

Score:

LAB #5: Shielding: Lead Aprons

Materials needed: Ion chamber, lead aprons.

DIRECTIONS:

Find out from your instructor what thickness of lead

is contained in the lead aprons you will be using . . . . . . . . . . . . : Lead Thickness:                  mm

Warm up the machine (3 exposures, 200 mA, 1 sec., 60 kVp, 5 seconds apart).

Set the ion chamber mode to “DOSE” or “INTEGRATE.” If there is a manual “RESET” button, push it before each exposure is made to re-zero the meter. If the meter reads out in R, you must convert all of the exposures to mR - to do this, move the decimal point on the meter to the right three spaces before listing it. NOTE: Many meters have a minimum sensitivity of 1 mR - therefore, any reading of zero should be listed below as “  1 mR” rather than as “0.” If the meter you are using reads in microroentgen, report the reading as measured.

EXPOSURE #1:  PATIENT IN-BEAM SKIN EXPOSURE:		DOSE: ___________

Place the meter on top of a 4" (10 cm) sponge			   (In mR)

and center the chamber area to the beam.

Collimate to 10" X 10" (25 X 25 cm) at

40" (100 cm) SID. Technique:

80 kVp, 150 mA @ 0.7 sec. = 105 mAs.

EXPOSURE #2:  IN-BEAM SKIN EXPOSURE WITH APRON ON:	DOSE: ___________

Repeat #1, but place a lead apron directly 			   (In mR)

over the meter.  Use the same technique.

PERCENTAGE OF

IN-BEAM DOSE:_______

(Divide dose #2 by dose

#1 and multiply by 100)

EXPOSURE #3:  IN-BEAM EXPOSURE WITH DOUBLE APRON:	DOSE:____________

Repeat #2 with two aprons draped directly

over the meter.  Use the same technique. 	    	PERCENTAGE OF

**BE SURE THE 2ND APRON IS HAS EQUAL LEAD**  	  IN-BEAM DOSE:________                   						 		   (Divide dose #3 by dose #1

and multiply by 100)

PERCENTAGE OF

SINGLE-APRON DOSE:               				    					   (Divide dose #3 by dose #2

and multiply by 100)

EXPOSURE #4: HIGH-KVP EXPOSURE WITH APRON ON:		DOSE:____________

Repeat #1 with one lead apron directly 			               (In mR)

over the meter, but use 110 kVp, 105 mAs.

PERCENTAGE OF

IN-BEAM DOSE:_______

(Divide dose #4 by dose

#1 and multiply by 100)

OPTIONAL - EXPOSURE #5: EFFECTIVENESS OF COMMON EYEGLASSES:

Place a couple of pairs of eyeglasses from classmates directly on a CR plate or cassette and expose using 40 kVp and 1 mAs.  Process the image and observe.

Questions:

1.  At 80 kVp, how effective was one lead apron at reducing exposure even in the direct beam of

x-rays?  Quantify by what percentage or ratio the primary beam was attenuated by the

apron:

2.  What does your answer to #1 mean personally to you about the general effectiveness of lead 	aprons in practice?

3.  How effective was a double apron at reducing exposure?  Quantify by what percentage or 	ratio the primary beam was attenuated by two aprons:

4.  Compare the first percentage in question #3 to the percentage in question #2. Does wearing a double apron absorb twice the radiation?  Will a radiographer’s exposure be inversely proportional to the number of aprons worn?

5.  Now compare the second percentage in question #3, the percentage of radiation transmitted

through a double-apron, to that transmitted through a single apron, (see Exp. #2). Was the

second apron more effective or less effective than the first? EXPLAIN why:

6.  Now look at the results for Exposure #4.  How effective was the lead apron at 	reducing

exposure at a high 110 kVp,?  Quantify by what percentage or ratio the primary beam

was attenuated by the apron:

7.  These figures are valid for shielding the patient from the primary beam.  Recall that aprons

are worn by radiographers to protect them from scatter radiation.  Had this experiment

been done with scattered radiation rather than the primary beam, how would you expect

the percentages of attenuation by the aprons to be different?  Why?

OPTIONAL #8.  Observe the radiograph of the eyeglasses.  Bear in mind that these were exposed to the primary beam. Remember that scatter radiation is usually of very low kVp as well as low intensity, much less than even this technique. Are eyeglasses effective in attenuating scatter radiation?  Was there was a significant difference between the two pairs of glasses?  If so, what might be a reason why?