Computerized Radiography (CR) 			Name:                                                            .

Score:

LAB #6: Sensitivity of CR Plates and Processing to Scatter and Background Radiation

Materials needed: Tissue-equivalent abdomen phantom (large dense phantom to generate scatter), CR cassette vertical holder.

Procedure

PART A: Sensitivity of CR Plates:

This part can be taken with a single exposure if three CR cassettes are available. If not, it can be done with one or two cassettes by simply repeating the exposures to the abdomen phantom and erasing the cassette between each exposure after processing.

Take a CR cassette from the bin and process it without erasing it. This exposure will be designated #0.

Place an abdomen phantom in the center of the x-ray table and set a technique of 80 kVp at 30 mAs. Be sure to turn off the AEC and use “manual techniques.” Erase three CR cassettes in the CR processor. Tape lead markers to them to designate them cassettes #1, #2 and #3. Using clamps, sponges or other means of support, place the first and second cassettes in vertical positions facing the abdomen - Place cassette #1 one meter away from the abdomen (at the end of the table), and cassette #2 two meters away on a cart, desk or chair. Keep cassette #3 behind a leaded wall outside of the room. Expose the abdomen phantom and process all three cassettes.

With the instructor’s permission, print out the images for side-by-side comparison on a viewbox.

Designate as #0 the printed film from the cassette that was taken directly from the bin and processed without erasing it.

PART B: Inability of Digital Processing to Compensate for a Pre-Fogged CR Plate:

Position an abdomen phantom in the center of the x-ray table (with no cassette plate under it). Center a 14 X 17” (35 X 43 cm) field to the abdomen, and set a technique of 110 kVp at 25 mAs to generate scatter radiation. Be sure to turn off the AEC and use “manual techniques.”

Erase a CR cassettes in the CR processor, and using a vertical cassette holder, sponges or other means of support, place it in vertical position facing the abdomen at a distance of one meter away from the abdomen (at the end of the table). This will be “Exposure A,” (a lead marker “A” can be placed on it). Expose the abdomen phantom and process the image of the scatter radiation for reference.

For Exposure B, repeat the above procedure to pre-fog the imaging plate, then USE THIS PRE-FOGGED PLATE placed under the abdomen phantom to take an abdomen radiograph. For the abdomen exposure, use 76 kVp and 20 mAs. Process the image.

For Exposure C, take a freshly-erased CR cassette and take a regular exposure of the abdomen phantom using 76 kVp and 20 mAs for comparison – this image should NOT be pre-fogged.

With the instructor’s permission, print out the images for side-by-side comparison on a viewbox.

Analysis: PART A

1. Compare image #1 with image #3 from the fully protected cassette. Is a fog density visible from cassette #1 placed at one meter from the abdomen?

2. Compare image #2 with image #3 from the fully protected cassette. Is a fog density visible

from cassette #2 placed at two meters from the abdomen?

3. Based on these results, what conclusion can you make regarding how much care must be taken

in protecting CR cassettes from scatter radiation?

4. Compare the unerased image #0 with image #3. Is a fog density visible from cassette #0 due to accumulated background radiation since the cassette was last erased?

5. Based on these results, what recommendation can you make to ensure that natural background radiation does not become a problem in degrading our CR images?

# Analysis: PART B

1. Observe Exposure A: Is there obvious fog present on the receptor plate? If so, this confirms that the condition of the plate for Exposure B is a pre-fogged condition. (If not, move the cassette closer to the phantom and repeat the entire series using this shorter distance.)

2. Hang the three images side-by-side for comparison; Note that Image B is the result of combining the fog apparent in Exposure A with the routine abdomen exposure apparent in Image C. Comparing images B and C, has the pre-fogged condition of the CR phosphor plate affected the final image of the abdomen phantom?

3. Was digital processing able to compensate for a pre-fogged condition of the CR phosphor

plate?

4. Would this limitation apply to DR  system which does not use a “chemical” plate?