Radiography Concepts	 			Name: _______________                     __

Score:

LAB #3: Focal Spot Size and Sharpness

Materials needed: A resolution test template; A small, dry bone, such as a phalanx; A smooth-edged coin, and several flat, stackable sponges to total at least 6 inches of thickness.

Procedure

Place a 10" X 12" (25 X 30 cm) CR cassette on the tabletop, centered to the x-ray beam. Collimate the field 1 inch (2.5 cm) smaller than the cassette on all sides, (and do not change  collimation throughout the experiment). Then stack sponges to create an OID of about 6 inches

(15 cm). Place a small dry bone and a smooth-edged coin very close to the central ray on top of the sponges. Place the resolution template alongside the coin and bone, crosswise to the table so that the lines in it are lengthwise with the table.

Be sure to turn off the AEC and use “manual” technique only. Make two exposures at 40" (100 cm) SID with the small and large focal spots as listed below. If possible, take both exposures using the same mA station. If not, adjust the exposure time to obtain equal mAs. With the instructor’s permission print out the images for side-by-side comparison.

Set a fixed kVp to 46 kVp

Exposure #1: 	Small focal spot,  2 mAs

Exposure #2:	Large focal spot, 2 mAs

Analysis

1. Observe the two images of the bone, noting the edges of small details such as the trabeculae in

the bone marrow. Can you tell visually which focal spot resolved higher sharpness of

detail? If so, which?

2. On each image of the resolution template, begin with the thickest set of lines and scan

downward through the sets until you first lose resolution of the lines. Have a partner

check to see if they agree where this point is. [Note that if you go beyond the first

complete blurring of the lines, they may seem to re-appear - this is false resolution, in

which each line blurs so much that it overlaps the next line. You must use the first blur

point.] Go to the previous set of lines - the last set in which the lines can be clearly

distinguished - and read the line-pair per millimeter number to the side. This is the

resolution. Record below:

Exposure #1: Small focal spot sharpness   =              LP/mm

Exposure #2: Large focal spot sharpness   =              LP/mm

3. Which SID produced the sharpest image? Explain why (See textbook):

4. As focal spot size increases, what happens to the sharpness of detail?

5. What is a good rule in practice for keeping sharpness at a maximum?

6. Compare the size of the two coin images. These may be overlapped for comparison. (Or, they

may be carefully measured. Magnification ratios may be calculated by dividing the real

coin width into each of the image widths.) Is there a significant difference in

magnification between the two focal spot sizes used?

(**NOTE that many digital systems can print out two images with somewhat different

sizes - Check the size of the entire image border to see if this has happened. If so, this

question must be answered with the coin image size relative to the overall image size.)

7. Review the textbook diagrams on focal spot size. Do both the umbra and the penumbra expand

with larger FS size? Is focal spot size a controlling factor for magnification of the gross

image?

8. Compare the shape of the two coin images. Is there any difference in distortion between the

two focal spot sizes used?

9. Examine the general brightness (density) and contrast of the small bone images. Does focal

spot size affect brightness or contrast in the image?

(**NOTE that many digital systems can print out two images with somewhat different

background densities - Check the overall brightness of the entire image background to 	see if this has happened. If so, this question must be answered with the bone density 	relative to the background density.)