Demonstrating Basic Digital Concepts			Name:_____________________________

Score:

Lab #4 C-Arm (Dose as a Function of Mode of Operation)

-By Bob Grossman

Materials needed: Tissue-equivalent abdomen, or lumbar spine, or skull, or hip phantom

Procedure

Configure the C-arm so that autocontrast is on and automatic brightness control is on.

For each set of conditions listed in the table below, do the following steps:

    - Record the OID.

    - Record the Pulse Rate if applicable.

    - Record the Initial Dose.

    - Take 10 exposures.

    - Record the Final Dose.

    - Save the 10th exposure and record the kVp and mAs.

    - Calculate the Dose/Image.

Dose/Image = (Final Dose – Inital Dose)/10

| Condition |  | Initial Dose | Final Dose | Dose/Image | kVp | mAs |
| --- | --- | --- | --- | --- | --- | --- |
| Normal Mode | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| Low Dose Mode | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| High Level Fluoro Mode | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| Digital Spot | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| High Pulse Rate Rate = ___ | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| Low Pulse Rate Rate = ___ | No Magnification No Collimation OID = __________ |  |  |  |  |  |
| Tight Collimation | Normal Mode No Magnification OID = __________ |  |  |  |  |  |
| Mag 1 | Normal Mode No Collimation OID = __________ |  |  |  |  |  |
| Mag 2 | Normal Mode No Collimation OID = __________ |  |  |  |  |  |
| Large OID OID = ____ | Normal Mode No Magnification No Collimation |  |  |  |  |  |
| Small OID OID = ____ | Normal Mode No Magnification No Collimation |  |  |  |  |  |

Analysis

- Rank the conditions from the most Dose/Image to the least.

- Rank the conditions from the best image quality to the least.

- How did the kVp and mAs change from Normal Mode to Low Dose Mode?

- How did the kVp and mAs change from Normal Mode to HLF Mode?

- How did the kVp and mAs change from Normal Mode to Digital Spot Mode?

- How did the kVp and mAs change from Normal Mode to the High and Low Pulse Rate conditions?

- Using the Normal Mode, High Pulse Rate, and Low Pulse Rate conditions, describe the relationship between pulse rate and dose.

- How did your kVp and mAs change when you increased your collimation?

- Describe the relationship between dose and collimation.

- Using the No Magnification, Mag 1, and Mag 2 conditions, describe the relationship between Dose and Digital Magnification.

- How did the kVp and mAs change when you increased digital magnification?

- Using the Small, Normal Mode, and Large OID conditions,  describe the relationship between Dose and Geometric Magnification.

- How did the kVp and mAs change when you increased geometric magnification?

Optional Additional Analysis

If the C-Arm displays brightness and contrast values you can record these values and describe how the brightness and contrast vary with the different conditions.

You can take a cine run that takes 10 seconds of fluoro time on the c-arm and calculate dose/frame for each condition.  You can also describe how motion appears differently when in pulse mode versus normal mode and at the different pulse rates.

You can take a cine run that takes 10 seconds as measured by a stop watch and compare that with the elapsed fluoro time when using the High and Low pulse rates.

Measure ROI values of a high attenuator and a low attenuator.    If your low attenuator has a mean value higher than your high attenuator, then your ROI values are reporting the level of darkness.  If your high attenator ROI mean value is higher than your low attenuator mean value, then your ROI values are already reporting the level of brightness and no conversion is required.

To convert level of darkness into brightness values on a 12-bit system convert each ROI value into a brightness value using the formula:

Brightness Value = 212  –  ROI Mean Value = 4096 – ROI Mean Value

*For a 14-bit system:  Brightness Value = 214  –  ROI Mean Value = 16384 – ROI Mean Value

Calculate contrast as the ratio of the High Attenuator Brightness Value / Low Attenuator Brightness Value.

Contrast = 	High Attenuator Brightness Value

Low Attenuator Brightness Value

Graph or describe how contrast and brightness changes with the different conditions.