Computerized Radiography (CR) 			Name:______________________________

Score:

LAB #2: Resolution: Film/Screen Vs. CR

Materials needed: Screen/Film Processing must be available for this lab; Step-wedge 				penetrometer, resolution test template, flat stackable sponges.

Procedure

Place a conventional screen cassette loaded with film on the tabletop. (List the type of screen used in the table below.) Stack flat sponges on top of the cassette to create a 6-8 inch OID. Place the resolution template on the sponges crosswise relative to the table, (so the actual pattern lines run lengthwise with the table). Place the step-wedge next to it so both are included within a single exposure. Use the large focal spot for all exposures. Be sure to turn of the AEC and use “manual technique.”

Recommended starting technique:  40 kVp at 2 mAs

Take a radiograph of the objects and process the film. Examine the resolution template for good density level. You may need to repeat the exposure to achieve a good medium density for the resolution template - it must not be too light or too dark.

Repeat the procedure using a CR plate placed on the tabletop and the same technique. Process and print out the CR image so the hard copies can be hung side-by-side on a viewbox for comparison.

Analysis

1. On the images of the resolution templates, scan visually from the larger bars to the thinner

ones and determine the first point at which they blur together. Read off the line-pairs per millimeter indicated on the template, and list below:

Type/Speed of Screen/Film Combination:

Screen/Film LP/mm  =   _____________

CR LP/mm   =   _____________

2. Which system produced the best spatial resolution:

3. On each the images of the step-wedge penetrometer, find two adjacent steps that demonstrate

medium-gray densities, not too light nor too dark. Measure each one using a

densitometer. List these in the formula below, and divide the lighter density value into the

darker one to obtain an indication of contrast resolution:

Screen/Film Contrast:       Darker      =                         =                     Contrast

Lighter

CR Contrast:       Darker      =                         =                     Contrast

Lighter

4. Which system produces higher contrast for the same two adjacent steps?

5. Which system has the advantage in spatial resolution (sharpness), and which system has the

advantage in contrast resolution?