Computerized Radiography (CR) 			Name:_____________________________

Score:

LAB #4: Effects of Technique on the Histogram

Materials needed: Tissue-equivalent knee phantom.

NOTE:  If histograms were traced from Lab #3, fill in the techniques used under A and B in the table at the top of page 2 and used these sketches to answer Questions #1 - 4. (Exposures C and D will need to be made at low and high kVp as indicated in the table.)

Procedure

Take CR radiographs of a knee phantom in AP projection as follows: Use the CR cassette plate  tabletop. Do not alter collimation, centering, or distances between views.

DO NOT DELIVER ANY IMAGES TO THE PACS OR PRINTER. You will need to copy the histogram and characteristic curve from the monitor. Lay a blank piece of paper over the monitor and carefully trace the graph directly: Trace the small square frame around the graph, the “bell-curve” graph portion, and the characteristic curve (the “S” curve).

Suggested starting “test” technique: 70 kVp at 4 mAs

Be sure to turn off the AEC and use “manual technique”

1. With the assistance of the instructor (or starting with the suggested technique above), establish a technique for the knee that produces an exposure indicator close to ideal (within 5%) as indicated by the equipment manufacturer. This may take a few repeated exposures. If you know this desired mAs in advance, no exposure for this technique will be required.

Kodak:   Target: 2000		Range: 	 1900 - 2100

Fuji:	   Target:   200 	Range: 	   190 - 210

Agfa:	   Target:    2.5   	Range: 	2.375 - 2.625

Phillips: Target:    110 	Range:		104.5 - 115.5

Write below this “ideal” technique.

“Ideal” Technique:                          mAs at 70 kVp

Take the following exposures; Trace and label the histograms for each:

Change in			Write In Actual

Technique			Technique Used

A. 	 1/4 mAs @ 70 kVp:		_____________

B. 	 4 X mAs @ 70 kVp:		 ____________

C. 	 Original mAs, 40 kVp:	_____________

D. 	 Original mAs, 100 kVp:	_____________

Analysis

Compare each of the histograms you traced side-by-side; Using a viewbox to shine through both sheets of paper, you can overlap the graphs for better comparison. Be sure to exactly line up the outside square border to compare shift of the curve. Line up the beginning point of the curves to compare their shape, and line up the two S-shaped characteristic curves to compare their slope.

Questions:

PART A: Effect of mAs

Compare the high-mAs histogram to the low-mAs histogram:

1. Did the high-mAs histogram shift in position right or left?

2. Why (or why not)?

3. Was this histogram curve constructed before or after normalization of the acquired

image? (See textbook).

4. Did the high-mAs histogram significantly change in its basic shape (change the relative

location of its peaks or number of peaks)?

5. Why (or why not)? What determines the number and relative location of peaks or

spikes in the histogram? (See textbook).

6. Did the characteristic curve significantly change in its slope; If so, steeper or

shallower?

7. Why ( or why not)?

PART B: Effect of kVp

Compare the high-kVp histogram to the low-kVp histogram:

8. Did the high-kVp histogram shift in position right or left?

9. Why (or why not)? Was this histogram curve constructed before or after normalization

of the acquired image? (See textbook).

10. Did the high-kVp histogram significantly change in shape (change the relative location of its peaks or number of peaks)?

11. Why (or why not)? What determines the number and relative location of peaks or

spikes in the histogram? (See textbook).

12. Did the characteristic curve significantly change in its slope; If so, steeper or

shallower?

13. Why (or why not)? Was this characteristic curve constructed before or after

normalization of the acquired image? (See textbook).