Computerized Radiography (CR) 			Name:______________________________

Score:

LAB #3: Exposure Parameters Using mAs

NOTE: The exposures at 1/4 and 4 times the original mAs from this lab can be used to save time during Lab #4; Use the graph pages from Lab #4 to trace the histograms from the CR display monitor for these two exposures only.

Materials needed: Tissue-equivalent knee phantom

Procedure

Take CR radiographs of a knee phantom as follows: Be sure to turn of the AEC and use “manual technique.” Use 70 kVp with a 10 X 12" (25 X 30 cm)  CR cassette tabletop. Do not alter collimation, centering, distances or kVp between views - only the mAs should be changed.

To best see the effects of quantum mottle, print all of the images so they can be compared side-by-side on a viewbox. Be sure to format the printer such that the images are equal to or larger than the original 10" X 12" plate. Use the “annotation” feature on the CR reader (processor) to mark on each image the mAs and kVp used. (It may be advisable to process each plate as soon as it is exposed in order to keep track of the images. If you lose track of which exposures are which, use the index number read-outs as a guide.)

Suggested starting “test” technique: 70 kVp at 4 mAs

1. With the assistance of the instructor (or starting with the suggested technique above), establish a technique for the knee that produces an exposure indicator close to ideal (within 5%) as indicated by the equipment manufacturer. This may take a few repeated exposures:

Kodak:   Target: 2000		Range: 	 1900 - 2100

Fuji:	   Target:   200 	Range: 	   190 - 210

Agfa:	   Target:    2.5   	Range: 	2.375 - 2.625

Phillips: Target:    110 	Range:		104.5 - 115.5

Write below this “ideal” technique.

“Ideal” Technique:                          mAs at 70 kVp

2. Cut the mAs in half in successive steps until mottle becomes readily apparent in the image

when compared with the original:

Mottle apparent?

@ 1/2 mAs:	     yes / no

@ 1/4 mAs:	     yes / no

@ 1/8 mAs):	     yes / no

3. Now, from the original technique, increase the original mAs in successive steps as listed 	below until saturation occurs or any other obvious degradation of the image:

Saturation / Degradation of Image Apparent?

@ 4X mAs:	     yes / no

@ 8X mAs:	     yes / no

@ 16X mAs:	     yes / no

(optional: @ 12X mAs:	     yes / no (do this only if 16X mAs did show changes)

Analysis

1. At what level of exposure reduction did mottle become apparent in the image?

2. Remember that slight variations in brightness and contrast are normal for CR processed images. In this lab we are looking for visually significant changes. At what level of exposure increase did saturation or any other significant degradation of the image

become apparent, if any?

3. For digital images, in which direction, decreasing technique or increasing technique, do we

have more exposure latitude (margin for error)? Is this a small or a large difference?

4. How does this contribute to the problem of dose creep as discussed in the textbook?

5. Why should dose creep be of concern, if the images are turning out well?

6. Since digital images do not manifest dose creep, what should radiographers pay attention to that does give an indication of it?