Quality Control		 			Name:______________________________

Score:

LAB #2: mA Linearity Test

Materials needed: Ion chamber

Actual rates of exposure can be measured by a typical ion chamber set to the rate mode, but it can be difficult to accurately read off these measurements, as they fluctuate during the exposure. For this lab, we will assume that the accuracy of the exposure times used has been established. (This can be done by conducting the QC tests in Lab #9.) In this way, integrated total exposure read-outs can then be used to check both for mA linearity and reproducibility.

Technique for all exposures:  70 kVp and 0.5 seconds

Procedure

For each of the following mA stations, re-set the ion chamber and take a reading immediately after exposure. Record the ion chamber read-outs below. If the meter reads in R, convert the

measurements into mR by moving the decimal 3 places to the right:

Exposure #1:   50 mA:                 mR

Exposure #2: 100 mA:                mR

Exposure #3: 200 mA:                mR

Analysis

1. We will use the 100 mA station (Exposure #2) as a reference point. For the 50 and 200 mA

stations, calculate the predicted percentage of exposure relative to the 100 mA station

by dividing the set mA by 100 mA, and multiplying by 100:

50 mA Station:    50     =                X   100   =              %   predicted

100

200 mA Station:   200     =                X   100   =              %   predicted

100

2. Now, for the 50 and 200 mA stations, calculate the measured percentage of exposure relative

to the 100 mA station by dividing the mR measured (above) by the mR measured at 100 mA, and multiplying by 100:

50 mA Station:                 mR (at 50 mA)  	 =                X   100   =              %

mR (at 100 mA)

200 mA Station:                 mR (at 200 mA)    =                X   100   =              %

mR (at 100 mA)

3. Comparing the measured percentages in Question #2 with the predicted percentages in

Question #1, were the 50 and 200 mA stations too high, too low, or accurate in their

output when compared to the 100 mA station?

50 mA Station:

200 mA Station:

4. For the 50 and 200 mA stations, calculate the percentage by which they are out of calibration

by dividing the smaller percentage from Questions #1 and #2 by the larger percentage,

then multiplying by 100:

50 mA Station:                 %    =                X   100   =              %

%

200 mA Station:                 %    =                X   100   =              %

%

5. What is the acceptable limit of deviation for mA linearity? (See textbook.)

6. Do each of these percentages from Question #4 fall within the acceptable limit?

50 mA Station:

200 mA Station: