Computerized Radiography (CR) 			Name:______________________________

Score:

LAB #5: Effects of kVp on the CR Image

Materials needed: Tissue-equivalent knee phantom.

Procedure

Take CR radiographs of a knee phantom as follows: Be sure to turn off the AEC and use “manual technique.” Use approximately 5 mAs with the CR cassette placed on the tabletop. Do not alter collimation, centering, distances or mAs between views - only the kVp  should be changed. Print all of the images so they can be compared side-by-side on a viewbox. Be sure to format the printer such that the images are equal or larger than the original 10" X 12" plate. Use the annotation feature to mark on each image the mAs and kVp used. (It may be advisable to process each plate as soon as it is exposed in order to keep track of the images. If you lose track of which exposures are which, use the index number read-outs as a guide.)

Suggested starting “test” technique: 70 kVp at 5 mAs

1. With the assistance of the instructor (or starting with the suggested technique above), establish a technique for the knee that produces an exposure indicator close to ideal (within 5%) as indicated by the equipment manufacturer. This may take a few repeated exposures:

Kodak:   Target: 2000		Range: 	 1900 - 2100

Fuji:	   Target:   200 	Range: 	   190 - 210

Agfa:	   Target:    2.5   	Range: 	2.375 - 2.625

Phillips: Target:    110 	Range:		104.5 - 115.5

Write below this “ideal” technique.

“Ideal” Technique:                          mAs at 70 kVp

2. Cut the kVp by 15% in successive steps until mottle becomes readily apparent in the image

when compared with the original:

Actual

Mottle apparent?	         kVp Used

@ minus 15% kVp:	     yes / no		         ________

@ minus 30% kVp:	     yes / no		         ________

(optional: @ minus 45% kVp:	     yes / no (do this only if minus 30% did not show mottle)

3. Increase the kVp by 15% in successive steps until saturation occurs or any other obvious degradation of the image:

Actual

Saturation / Degradation of Image Apparent?	kVp Used

@ plus 15% kVp:	     yes / no					________

@ plus 30% kVp:	     yes / no					________

@ plus 45% kVp:	     yes / no					________

Analysis - Part A: Exposure Parameters

1. At what level of kVp reduction did mottle become apparent in the image?

2. Remember that slight variations in brightness and contrast are normal for CR processed images. In this lab we are looking for visually significant changes. At what level of kVp  increase did saturation or any other significant degradation of the image become apparent, if any?

3. Compare these images and results with those from Lab #3 in which the mAs was varied up and down: Do you come to the same conclusion as you did for Lab #3 regarding in which direction (decreasing or increasing technique) we are allowed more exposure latitude (margin for error)?

4. If kVp controls beam penetration, whereas mAs controls beam intensity, why do we get such

similar lab results taking into account their effects on the remnant x-ray beam?

5. Why is sufficient penetration still absolutely essential for digital imaging just as it has always been for conventional radiography?

Analysis - Part B: Image Quality

6. Remember that slight variations in brightness and contrast are normal for CR processed

images. In this lab we are looking for visually significant changes. Leave aside the two most extreme images and observe the remaining ones: As kVp is increased or decreased in one or two steps of 15%, do you visually see significant changes for those images in:

A. Brightness: 	yes / no

B. Contrast:		yes / no

7. With conventional screen/film radiography, substantial changes in both brightness and 	contrast occur with these types of changes in kVp. Explain why you got these answers to 	question #6 for CR radiography: