Quality Control		 			Name:______________________________

Score:

LAB #1: Timer Quality Control Tests

Materials needed: An electrical spinning top test device for a three-phase or high-frequency x-ray machine, and a protractor for measuring angles; A manual spinning top test device for a single-phase machine.

Procedure

A: FOR THREE-PHASE OR HIGH-FREQUENCY MACHINE:

Place the test device on a CR cassette and test the accuracy of the two exposure times

below, or select other exposure times less than 1/4 (0.25) seconds for this device.

Set a technique of 60 kVp and 50 mA.

Exposure #1: 0.1 (1/10) second

Exposure #2: 0.2 (1/5) second

Results:

1. For each arc of density produced on the radiographs, carefully center a protractor

(angligner) to the center axis point of the spinning top, then measure the degrees

of arc and record below:

Exposure #1:               degrees measured

Exposure #2:               degrees measured

2. Electrical spinning tops are normally designed to spin at precisely one revolution per second (1 rps), which is 360 degrees per second. Multiply the set exposure times by 360 to obtain the degrees of arc that should have been produced at these times:

Exposure #1:               degrees predicted

Exposure #2:               degrees predicted

B: FOR SINGLE-PHASE MACHINE:

Place the manual spinning top test device on a CR cassette and test the accuracy of the

two exposure times below, or select other exposure times less than 1/2 (0.5) seconds for

this device. For each test, the top must be manually spun and the exposure taken right

away before it slows down too much.

Set a technique of 60 kVp and 100 mA.

Exposure #1: 0.07 (1/15) second

Exposure #2: 0.15 (1/6) second

Results:

1. For each radiograph, simply count the number of density dots produced and record

below. (Note that the dots must be well separated; if they are not, the exposure

must be repeated with the top spun faster or the exposure made more quickly after

it is spun.) When counting the dots, you may include estimated fractions of a dot.

Exposure #1:               dots counted

Exposure #2:               dots counted

2. For a fully-rectified single-phase machine, 120 dots per second should be produced.

Simply multiply the set exposure times by 120 to obtain the number of dots that

should have been produced at these times:

Exposure #1:               dots predicted

Exposure #2:               dots predicted

Analysis

3. For each exposure, compare the degrees or dot count measured in Question #1 with the

degrees or dot count predicted in Question #2; Was the actually exposure time too long,

too short, or accurate?

Exposure #1:

Exposure #2:

4. For each exposure time tested, divide the smaller number between Question #1 and Question

#2 by the larger number, then multiply by 100 to obtain the percentage by which the

actual time is off:

Exposure #1:	  Smaller #     =                   =                %

Larger #

Exposure #2:	  Smaller #     =                   =                %

Larger #

5. What is the acceptable limit of accuracy for exposure timers? (See textbook.)

6. For each exposure tested, does the actual time fall within the acceptable limits of accuracy?

Exposure #1:

Exposure #2: