Radiation Protection					Name:______________________________

Score:

LAB #4: Half Value Layer: Filtration and kVp

Materials needed: Ion chamber; Various slabs of aluminum, recommended thicknesses: 1 mm,

0.5 mm, and 0.2 mm, several of each.

Half value layer is defined as the amount of absorber required to reduce the intensity of the x-ray beam to one-half the original intensity. It is the radiographic unit of measure for x-ray beam penetration. Part A of this lab examines the effect of added filtration upon beam penetration, Part B examines the effect of kVp upon penetration.

PART A: Effect of Added Filtration on HVL

OPTION #1: If the collimator can be loosened, or added filter slab accessed:

*Option #1 must be done under the immediate supervision of an instructor.

On many x-ray machines, simple thumb-screws where the top of the collimator attaches to the x-ray tube housing can be loosened to access a slab of added aluminum filter on top of the collimator box. (Be sure not to loosen the screws all the way, as the collimator may drop.) Loosened to leave about a 1/8" or 2 mm gap, the metal leg from a set of measuring calipers can be inserted to push the aluminum slab out the opposite side. Re-tighten the screws.

When replacing this slab, be sure to carefully center it from each side of the collimator

before tightening the screws. The leg of the calipers can be used to push the slab into

place, then re-tighten the screws.  Do not forget to replace the slab!

Other machines may have a slot in the upper collimator, or otherwise allow easy access to

the slab of added aluminum. In most cases, the slab is 1 mm thick. For Part A, compare

the HVL without the slab to the HVL with the slab added back in.

OPTION #2: If added filtration in or above the collimator cannot be accessed:

If no access can be gained to the added filtration, the concept of the lab can still be

demonstrated by using the x-ray tube as is for the “unfiltered beam,” then (after

collimation and centering are completed) tape 2 mm of additional aluminum filter slabs

on the bottom of the collimator. The collimated field should not extend beyond these

slabs. This will serve as the “filtered beam,” which will be compared with the “unfiltered

beam” for Part A.

Perform the entire procedure below first for an “unfiltered beam,” then for a “filtered” beam.

Directions:

1. Set the SID to 44" (110 cm). Set technique close to 75 kVp, 100 mA at 0.33 seconds, tabletop mode.

2. Place the ion chamber on the table centered to the x-ray beam. Collimate to a 4" X 4" (10 X 10

cm) field.

3. Make an exposure and record the reading below as the “no absorber”measurement. Note that 	if the meter reads in R, you will need to move the decimal 3 places to the right (multiply 	by 1000) to report the reading in mR.

4. Be sure to re-set the ion chamber to zero as needed before each exposure is made. Place a 1

mm slab of aluminum directly on the ion chamber, repeat the exposure, and log the

measurement below. Continue adding 1 mm slabs until the measurement reads less than

one-half the original, “no absorber” measurement.

5. Remove the last slab of aluminum and replace it with an 0.5 mm slab. Record the measured

exposure and the amount of aluminum absorber. If this measurement is more than one-

half the original, add 0.1 mm slabs in sequence until the measurement is less than one-

half. If it is already less than one-half, remove the 0.5 mm slab and then add 0.1 mm slabs

in sequence until the measurement is less than one-half the original.

6. There should now be at least 5 or 6 measurements, which can be plotted on the accompanying

graph.

7. Repeat the above procedure for a “filtered beam,” with the original aluminum slab added back

into the collimator (for Option #1) or with 2 mm of aluminum taped to the bottom of the

collimator (for Option #2). You will find that you can start with a higher amount of

absorber and use larger increments on this second series, to save time.

8. From the tables you have filled out, carefully plot the measurements taken for each series on

the appropriate graph attached.

9. Using a pencil to start with, construct a smoothly curving “best-fit” line through the data points

plotted on each graph. NOTE that often a measurement or two will sharply differ from 	the general trend, falling above or below the expected line. For these, you must “split the

difference” between this point and the adjacent 2 points so that a smoothly-curving and

continuous line is developed. It will likely take several tries. This is the most accurate 	line taking all measurements into consideration. Draw the final line in ink.

10. For each plotted graph, find the exposure level which is precisely one-half of the original “no

absorber” reading. Using a ruler, carefully draw a straight, precisely horizontal line from

this point until it intersects the curve you drew. Then drop a straight, precisely vertical

line from that point down to the bottom axis of the graph. This amount of absorber is the

HVL.

11. Log the HVL for each x-ray beam on the last page - Specify the type of absorber used with

your units. Answer the questions for Part A.

[Scoring suggestions for the Instructor: Points may be deducted for 1) inaccurate plotting of points, 2) a kinked or uneven “best-fit” line, 3) failure to draw in lines for accurate reading, 4) reading lines that are not accurately horizontal and vertical, or not aligned properly. The students’ results should fall within + or - 0.2 mm of the instructor’s reading for the “no absorber” series, and within + or - 0.3 for the other two series, which have more condensed measurements.]

PART B: Effect of Increased kVp on HVL

Repeat the entire procedure from Part A for an x-ray beam with 120 kVp. Use the same mA and exposure time settings, and identical collimation, distance and centering. Record your measurements in the table for “High kVp,” plot the points on the “High kVp” graph, and carefully construct a “best-fit” line on this graph. Use carefully drawn horizontal and vertical lines to interpret the graph, and report the HVL on the last page with proper units and specifying the type of absorber. Answer the questions for Part B.

Measurement Tables:

Unfiltered Beam

70 kVp

# Absorber  Intensity (mR)

No absorber   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

Filtered Beam

70 kVp

# Absorber  Intensity (mR)

No absorber   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

High kVp Beam

120 kVp

# Absorber  Intensity (mR)

No absorber   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

__________   __________

Questions:

From your graphs, list the measured HVL for each beam below, (Be sure to use correct units):

Without Filtration, 70 kVp: _________________________________________

With Filtration, 70 kVp: _________________________________________

With Filtration, 120 kVp: _________________________________________

PART A:   Compare the two 70-kVp measurements:

1. Which one has a higher HVL?

2. What does this mean about the relationship between filtration and x-ray penetration?

3. What does this imply about the relationship between filtration and the amount of

radiation that gets absorbed within the patient?

PART B:   Compare the two filtered measurements:

1. Which one has a higher HVL?

2. What does this mean about the relationship between kVp and x-ray penetration?

3. What does this imply about the relationship between kVp and the percentage of

radiation that gets absorbed within the patient?