Radiation Protection					Name: _____________________________

Score:

LAB #6: Exposure Levels and Barriers

Materials needed: Ion chamber, plexiglass abdomen phantom.

DIRECTIONS:

Warm up the machine (3 exposures, 200 mA, 1 sec., 60 kVp, 5 seconds apart).

Set the ion chamber mode to “DOSE” or “INTEGRATE.” If there is a manual “RESET” button, push it before each exposure is made to re-zero the meter. If the meter reads out in R, you must convert all of the exposures to mR - to do this, move the decimal point on the meter to the right three spaces before listing it. NOTE: Many meters have a minimum sensitivity of 1 mR - therefore, any reading of zero should be listed below as “  1 mR” rather than as “0.” If the meter you are using reads in microroentgen, report the reading as measured.

PART A:

EXPOSURE #1:  PATIENT IN-BEAM SKIN EXPOSURE:		DOSE: ___________

Place the meter on top of a 4" (10 cm) sponge			   (In mR)

and center the black area to the beam.

Collimate to 10" X 10" (25 X 25 cm) at 40"

(100 cm) SID. Technique:

100 kVp, 150 mA @ 0.7 sec. = 105 mAs.

EXPOSURE #2:  PT. SCATTER EXPOSURE OUTSIDE BEAM:		DOSE: ___________

Place the ABDOMEN plexiglass phantom in 			   (In mR)

the light field, centering so that the edge of

the 10" X 10" field is right along the top of  	 PERCENTAGE OF

the phantom.  This creates a realistic source  	    IN-BEAM DOSE:_______

of scatter.  Place the meter on its side facing 	      (Divide this dose by the

the top of the phantom and slide it up 1 inch  	       dose in #1 and multiply

from the top edge of the light field (and 		        by 100.)

phantom).  Use the same technique as in #1.

EXPOSURE #3:  OPERATOR EXPOSURE AT ONE METER:		DOSE: ___________

Keep the phantom in the beam as in #2.  Place

the meter on its side facing the top of the	 	PERCENTAGE OF

phantom, but slide it to 1 meter from the 		  IN-BEAM DOSE:_______

top edge of the light field and phantom.  		    (Divide this dose by the

Use the same technique as in #1.			     dose in #1, and X 100)

Questions:

1.  At this technique, how much difference did it make whether a particular organ lay within the

direct x-ray beam or just outside the edge of the beam?  (Quantify the ratio or percentage):

2.  What does your answer to #1 imply that you should do in practice?

3.  At this technique, how does the technologist’s exposure standing 1 meter away compare with

the patient’s in-beam exposure?  (Quantify the ratio or percentage to the nearest 10th):

4.  At this technique, how much more scatter radiation does the technologist receive if she stands 	right against the patient as compared to standing 1 meter away?  Find the ratio of 		difference by simply dividing the dose in #3 into the dose in #2 - Quantify how many

“times more” radiation the tech recieves:

5.  What does your answer to #4 imply that you should do in practice?

PART B: **NOTE: This portion of the lab is OPTIONAL and should only be performed under the direct supervision of an instructor.

EXPOSURE #4:  DIRECT EXPOSURE OUTSIDE 		      DOSE: _____________

CONTROL BOOTH:						(In mR)

Aim the tube horizontally directly at the

control booth window and 1 meter away.

Place the ion chamber on the window

ledge outside the booth, on its edge facing the

tube.  Collimate to the size of the meter.  Use

100 kVp, 150 mA @ 0.7 sec. = 105 mAs.

*Stand to the side of the window, but

behind the booth wall, while shooting.

EXPOSURE #5:  DIRECT EXPOSURE INSIDE			     DOSE: _____________

CONTROL BOOTH WINDOW:					(In mR)

Repeat the procedure in #4 with the same

technique, only place the meter on its side         PERCENTAGE OF

on the window ledge inside the control 	   	IN-BEAM DOSE:________

booth, facing the tube.				(Divide this dose by the dose

*Stand to the side of the window, but 	     	  in #4 above, then X 100)

behind the booth wall, while shooting.

Questions:

6.  At this technique, how effective was the leaded window of the control booth at attenuating

x-rays from the primary beam?  (Quantify the percentage or ratio):

7.  What does your answer to #6 mean to you in practice, in terms of how safe the control booth 	is?