Radiation Protection					Name:______________________________

Score:

# LAB #8: Scatter Exposure to Personnel Vs. Location

by Dennis Bowman, R.T.

Materials needed: Ion chamber, CR plates, hand phantom, devices to securely hold the CR plate

and ion chamber in a vertical position.

DIRECTIONS:

Warm up the machine (3 exposures, 200 mA, 1 sec., 60 kVp, 5 seconds apart).

Set the ion chamber mode to “DOSE” or “INTEGRATE.” If there is a manual “RESET” button, push it before each exposure is made to re-zero the meter. If the meter reads out in R, you must convert all of the exposures to mR - to do this, move the decimal point on the meter to the right three spaces before listing it. NOTE: Many meters have a minimum sensitivity of 1 mR - therefore, any reading of zero should be listed below as “  1 mR” rather than as “0.” If the meter you are using reads in microroentgen, report the reading as measured.

Place an abdomen phantom upright on a gurney (stretcher) that is raised to the sitting position. Place a 14" X 17" (35 X 42 cm) image receptor plate behind the phantom. Angle the x-ray tube downward at 15 degrees, with an SID of 72" (180 cm), and centered to the phantom with the field opened to fill the IR plate.

Set an appropriate technique for a 72" (180 cm) abdomen exposure, (recommended 80 kVp at 80 mAs). Use the same technique for all exposures, and do not change distance, collimation or centering.

At each of the angles listed below, the ion chamber will be placed at distances of 3 feet, 6 feet and 9 feet (1, 2 and 3 meters) from the phantom.

At selected positions you wish to compare, also place a CR plate with the hand phantom directly in front of it. Position it so it is firmly supported in a vertical position facing the phantom, next to the ion chamber. (You will not want to do this for every single position - choose those positions of most interest to you.) Use lead markers taped to the IR plate to keep track of which positions these are placed in.

For each exposure, note below the read-out of the ion chamber. Also process the hand images.

SERIES A: 90 degrees from the central ray, directly to the side of the “patient”:

Exposure #1: At 3 feet (1 meter):			DOSE: ____________

(In mR)

Exposure #2: At 6 feet (2 meters):			DOSE: ____________

(In mR)

Exposure #1: At 9 feet (3 meters):			DOSE: ____________

(In mR)

SERIES B: 45 degrees forward from the “patient”:

Exposure #1: At 3 feet (1 meter):			DOSE: ____________

(In mR)

Exposure #2: At 6 feet (2 meters):			DOSE: ____________

(In mR)

Exposure #1: At 9 feet (3 meters):			DOSE: ____________

(In mR)

SERIES C: 45 degrees behind the “patient”:

Exposure #1: At 3 feet (1 meter):			DOSE: ____________

(In mR)

Exposure #2: At 6 feet (2 meters):			DOSE: ____________

(In mR)

Exposure #1: At 9 feet (3 meters):			DOSE: ____________

(In mR)

SERIES D (OPTIONAL): 0 degrees, directly behind the “patient.” NOTE that the detectors 			must be positioned so that they are a least a few inches above the estimated 			edge of the remnant x-ray beam, so they are measuring ONLY scatter.

Exposure #1: At 3 feet (1 meter):			DOSE: ____________

(In mR)

Exposure #2: At 6 feet (2 meters):			DOSE: ____________

(In mR)

Exposure #1: At 9 feet (3 meters):			DOSE: ____________

(In mR)

SERIES E (OPTIONAL): 180 degrees, directly in front of the “patient.” NOTE that the 			detectors must be positioned so that they are at least a few inches below the 			estimated edge of the primary x-ray beam, so they are measuring ONLY 			scatter radiation.

Exposure #1: At 3 feet (1 meter):			DOSE: ____________

(In mR)

Exposure #2: At 6 feet (2 meters):			DOSE: ____________

(In mR)

Exposure #1: At 9 feet (3 meters):			DOSE: ____________

(In mR)

Questions:

1. Rank below the angles from which the highest to the lowest scatter exposure was received to

personnel from the “patient”:

1.

2.

3.

4.

5.

2. Based on your observations, where is the safest place for radiographers to stand, relative

to the patient, when they must be in the room during an exposure?

3. Based on your observations, where is the worst place for radiographers to stand, relative

to the patient, when they must be in the room during an exposure?

4. Choose the series for which you obtained the highest exposure readings. Divide the read-out

for 6 feet (2 meters) into the read-out for 3 feet (1 meter); Fill in the result below:

1 meter reading     =                 mR   =                  fraction of original exposure

2 meter reading                        mR

According to the inverse square law, the 2 meter reading should be approximately 1/4

(0.25) of the 1 meter reading. Is the result above close to this amount? Does

scatter radiation follow the inverse square law?

5. Hang all of the CR images of the hand on viewbox illuminators for comparison. What your

impressions of the scatter radiation levels in the x-ray room? What are your impressions

regarding the sensitivity of CR imaging plates? Do these images confirm the read-outs

from the ion chamber?