Computerized Radiography (CR) 			Name:______________________________

Score:

LAB #8: Attempts to Corrupt the Exposure Indicator

-by Dennis Bowman

Materials needed: Tissue-equivalent elbow (or similar) phantom.

Procedure

Take all exposures tabletop. Be sure to turn off the AEC and use “manual technique.”

To calculate the percentage by which each exposure indicator read-out is changed from the first  properly-exposed image, use the formula:

O - N     X 100

N

where O is the original exposure indicator read-out and N is the new EI read-out.

A. Centering over different tissues: On a single 14" X 17" (35 X 42 cm) CR plate, take an exposure of an abdomen (or thorax) phantom centered and collimated correctly. Use a technique chart or consult your instructor for the best technique to use for a proper exposure. Exposure #1 must generate and exposure indicator read-out that falls within the “ideal” range recommended by the manufacturer. After processing the image, note the value of the exposure indicator below for “Exposure #1.” Then off-center the phantom 2" higher and 2" lower than normal, without changing collimation or any other variable. The x-ray beam remains centered to the image receptor plate. Note below the exposure indicators produced, and calculate for each (using the above formula) the percentage by which the EI was altered:

Exposure Indicator		Percentage

Readout         		   Change_

Exposure #1: Phantom properly centered: 	                              		________

Exposure #2: Phantom centered 2" high: 	                              		________

Exposure #3: Phantom centered 2" low: 	                              		________

B. Off-centering to include excess background: On a single 10" X 12" (25 X 30 cm) CR plate, take an exposure of an elbow or ankle phantom centered and collimated correctly. Use a technique chart or consult your instructor for the best technique to use for a proper exposure. Exposure #1 must generate and exposure indicator read-out that falls within the “ideal” range recommended by the manufacturer. After processing the image, note the value of the exposure indicator below. Then off-center the phantom as far as possible to one side without clipping it off. Do not change collimation or any other variable. Note below the exposure indicators produced, and calculate for each (using the above formula) the percentage by which the EI was altered:

Exposure Indicator		Percentage

Readout         		   Change_

Exposure #1: Phantom properly centered: 	                              		________

Exposure #2: Phantom centered 4" to side: 	                              		________

C. Centering over different tissues: On a single 14" X 17" CR plate, take an exposure of an knee  phantom properly centered, but collimated to a 6" X 8" field . Use a technique chart or consult your instructor for the best technique to use for a proper exposure. After processing the image, note the value of the exposure indicator below. For Exposure #2,  open the field so there are 2" of “raw” beam on each side of the knee phantom. For Exposure #3, open the field so there are about 5" of “raw” beam on each side. For Exposure #4, close the collimated field to only a 2" square in the middle of the phantom.

Note below the exposure indicators produced, and calculate for each (using the above formula) the percentage by which the EI was altered:

Exposure Indicator		Percentage

Readout         		   Change_

Exposure #1: Properly collimated:	 	                              		________

Exposure #2: Collimation open 2":	 	                              		________

Exposure #3: Collimation open 5":	 	                              		________

Exposure #4: Closed to 2" square field: 	                              		________

Process all exposures. Print them out only with the instructor’s permission.

Analysis

1. Observe your results for Part A; Did off-centering over different types of tissue corrupt the exposure indicator value by more than 5%?

2. Observe your results for Part B; Did off-centering such that 4" of “raw” beam was included

within the exposure field corrupt the exposure indicator value by more than 5%?

3. Observe your results for Part C:

A. Did opening the collimated field 2" all around corrupt the exposure indicator value by

more than 5%?

B. Did opening the collimated field 5" all around corrupt the exposure indicator value by

more than 5%?

C. Did closing the collimated field to a small 2" square corrupt the exposure indicator

value by more than 5%?

4. For 3C above, what rule for using CR plates directly applies to this effect? (See textbook.)