Quality Control		 			Name:____________________________

Score:

LAB #5: Beam Alignment and Field Size Accuracy

Materials needed: A ruler. Paper clips (unfolded into an “L” shape) or other radiopaque objects

can be used to mark the corners of the light field, however, more accurate

test templates are available and may include a plexiglass cylinder device

for checking the verticality of the central ray - these are recommended.

Procedure

The cassette plate used must be larger than the field size being checked. (Check an 8 X 10 inch field on a 10 X 12 inch cassette plate.) Use a 40-inch (100 cm) SID. With the instructor’s permission, print out your images. Be sure to turn off the AEC and use “manual technique.”

Technique for all exposures: 50 kVp and 2 mAs.

“Homemade” Method: If using the “home-made” method, carefully mark  with paper clips the

corners and mid-edges of the light field all around. Also, place a metal marker precisely centered to the cross-hairs of the light field. A marker must be placed in one corner of the field to indicate the orientation of the entire field - make a note whether this corner is northeast, northwest, southeast, or southwest.

Commercial Method: If a commercial device is used, first, carefully center the cross-hairs of the

light field should to the center point marker. Do not re-center in order

to align the edges of the field. Collimate the edges of the light field as

closely as possible to the marked borders on the plate all around. There

will be a marker placed in one corner of the field to indicate the 					orientation of the entire field - make a note whether this corner is

northeast, northwest, southeast, or southwest.

If a plexiglass cylinder is available to check for central ray verticality,

first check the angle adjustment on the x-ray tube housing and set it as

close to zero as possible. Place the cylinder on the test plate and center the

bottom washer or marker to the cross-hairs of the light field as seen on the

test plate.

Analysis

A. FIELD SIZE:

1. Measure the length and width of the actual x-ray field as the black area on the resulting image.

On the original device, measure the length and width of the line-out field area (or paper

clips). Record below (with proper units, cm or inches) and compare:

Length:   Setting:                               Measured: ________________

Width:    Setting:                               Measured: ________________

2. What is the acceptable limit of deviation for field size accuracy? (See textbook.)

3. Are these measurements within the limits of acceptability? (To find the percentage of

deviation for each, subtract the smaller number from the larger, then divide this

difference by the field size setting and multiply by 100.)

4. In correcting for any inaccuracies, should the x-ray field be adjusted to match the light field, or

should the light field be adjusted to accurately indicate the location of the x-ray field?

B. FIELD EDGES:

5. For each edge of the light field, measure how far the light field markers are from the edge of

the black exposed area on the image, and in which direction (inside the black area, or

outside). NOTE that it is the light field, (not the x-ray beam), which is being checked and

which would be corrected. The location of the light field is indicated by the markers.

List your results below:

North edge   =                    inches/cm off to the                             (inside or outside)

South edge   =                    inches/cm off to the                             (inside or outside)

East edge     =                    inches/cm off to the                             (inside or outside)

West edge    =                    inches/cm off to the                             (inside or outside)

C. CENTRAL RAY LOCATION:

6. Draw diagonal lines from each pair of corners of the black exposed area on the image in order

to locate its exact center point where these lines cross. Repeat this procedure using the

markers in the image to find the center of the light field. Measure with a ruler the 	distance between these two center points, note in which direction the center of the light 	field is off from the center of the x-ray field, and record below:

Light field center is                (inches/cm) off toward the                  (direction) from the

center of the x-ray field.

7. Divide this measurement from #6 above by the SID (using the same units), and multiply the

result by 100 to obtain the percentage of central ray deviation and record:

Shift     =                       X   100   =                     %

SID

8. From the textbook, what is the limit of acceptable deviation for the central ray of the light

field?

9. Is your calculation in #7 within this limit?

D. VERTICALITY OF CENTRAL RAY:

10. If a verticality cylinder was used, note the location of the image of the BB or screw that was

in the top of the cylinder. It should fall within the outer circle (about 1 cm from the

center-point) of the washer or leaded circles. Normally, two circles are provided, with the

inner circle (or inner edge of the washer) indicating 1 degree of deviation from vertical,

and the outer circle (or outer edge of the washer) indicating 2 degrees of deviation from

vertical.

Note that although the angle of the x-ray tube lengthwise to the table is adjustable and

therefore not of critical importance for this test, the crosswise angle of the x-ray tube is

usually locked into place and cannot be changed by the radiographer - This angle should

be perfectly vertical.

Use the corner marker on the image to orient the image.

In the crosswise direction to the x-ray table, is the central ray of the x-ray beam vertical

within plus or minus 2 degrees? If not, in which direction does it need to be corrected?